## Sheet1
Date:
Objective
Reorienting the magnetization direction of a ferromagnetic material
S.No.: | Procedure | Comment
1 | Switch on the chiller, and set it to 1deg C, 30 min prior to the experiment | Chiller takes around 30 min to reach 1 deg C,
2 | Place the magnetic thermal annealing (MTA) setup in between the poles of the Electromagnet
3 | Open the sample holder tray of the MTA setup and mount your sample, put the sample holder tray back inside and screw it well
4 | Now adjust the MTA setup such that the heating chamber, inside which sample is there, is present in the middle of the poles | Constant magnetic field will be present only in between the poles
5 | Place two thermometers touching the poles, make sure they do not touch the heating chamber | For monitoring the poles temperature
6 | Connect the yellow and heating wire from MTA to the heater
7 | Switch on the heater and program its PID through 'Temperature' application available on our lab computer, program the PID for heating at 400 deg C for 1 Hour
8 | Once the chiller temperature reaches 1 deg C, switch on the pirrani gauge | To monitor the pressure generated at the vacuump pump
9 | Open the lab window and open the vacuum pump valve just outside
10 | Close the Right-angel valve at the top of MTA setup, start the vacuum pump, once the pirrani gauge reads in the order 10e-2, open the right angel valve
11 | Start the heater by pressing 'UP ARROW' physical button on the heater PID, then switch on the heater toggle button
12 | Wait until the heater reaches 400 deg C, in the meantime switch on the CCPS, connect the RS232 to the lab laptop, open the 'EMController' application
13 | When the heater reaches 400 deg C, select constant field in the measurement tab, then check the 'PWM' option and set it to 68% (this will give 400mT)
14 | Set the CCPS threshold temperature to 65 deg C | It is the coil temperature, it should be more than 65 deg C
15 | Through out the experiment keep an eye on the poles temperature, if it reaches 70 deg C in either of the pole, immeditaely stop the heater and increase the pole gap